# Biomedical ECG Signal Analysis using Python and Arduino

Technical Progress Report

Generated on: 08 May 2026

# 1. Introduction

This project focuses on the development of a real-time
Biomedical ECG Signal Analysis system using Python and Arduino.
The objective is to acquire ECG-related signals, process them
using Digital Signal Processing (DSP) techniques, and prepare
a Hardware-in-the-Loop (HIL) architecture for real-time
cardiac monitoring applications.

The project integrates concepts from:
- Signals and Systems
- Biomedical Signal Processing
- Embedded Systems
- Real-Time Data Acquisition
- Hardware-in-the-Loop (HIL) Systems

The final system architecture is designed to support:
- ECG signal acquisition
- Real-time filtering
- QRS complex detection
- Heart rate estimation
- Arrhythmia classification
- Live graphical visualization

# 2. Project Objectives

Acquire ECG signals using hardware and software systems

Implement digital filtering techniques

Detect QRS complexes in real time

Calculate heart rate from ECG signals

Prepare a Hardware-in-the-Loop (HIL) system architecture

Develop a real-time biomedical monitoring dashboard

Integrate Arduino communication with Python

# 3. Software Development Phase

The software development phase was completed prior to
hardware integration in order to stabilize the Digital
Signal Processing pipeline.

The ECG processing system was initially developed using
offline ECG datasets. The MIT-BIH Arrhythmia Database
was used as the primary ECG signal source.

The following stages were implemented:

ECG dataset acquisition

Signal visualization

Bandpass filtering (0.5 Hz – 40 Hz)

QRS peak detection

Heart rate calculation

Arrhythmia classification

Real-time ECG simulation

Real-time GUI dashboard creation

# 4. Digital Signal Processing

Digital Signal Processing (DSP) techniques were implemented
to remove noise and improve ECG signal quality.

A Butterworth bandpass filter was designed with:
- Lower cutoff frequency: 0.5 Hz
- Upper cutoff frequency: 40 Hz

The filtering stage removes:
- Baseline wander
- High-frequency noise
- Motion artifacts

The filtering was implemented using the scipy.signal library.

# 5. QRS Detection and Heart Rate Estimation

QRS complexes were detected using peak detection algorithms.

The scipy.signal.find_peaks() function was used to identify
R-peaks from the filtered ECG waveform.

Heart rate was calculated using RR intervals according to:

HR = 60 / RR_interval

where RR_interval represents the average time between
consecutive R-peaks.

# 6. Real-Time ECG Dashboard

A real-time ECG dashboard was developed using PyQtGraph.

The dashboard supports:
- Live ECG waveform visualization
- Real-time QRS detection
- Live BPM display
- Arrhythmia condition display
- Real-time updating architecture

A moving average technique was implemented to smooth
heart rate fluctuations and improve stability.

# 7. Hardware-in-the-Loop (HIL) Preparation

The project architecture was designed as a Hardware-in-the-Loop
(HIL) biomedical monitoring system.

The software system was structured to support both:
1. Offline ECG dataset processing
2. Real-time Arduino serial acquisition

The architecture separates:
- Data acquisition
- DSP processing
- Detection
- Visualization

This modular design simplifies future hardware integration.

# 8. Arduino-Python Communication

Before integrating the ECG sensor module, serial communication
between Arduino and Python was established and verified.

The Arduino Uno was programmed to transmit serial data through
USB communication using a baud rate of 115200.

Python communication was implemented using the PySerial library.

The following stages were completed:

Arduino serial transmission setup

COM port identification

Python serial communication

Real-time data reception in Python

Continuous streaming verification

Analog input streaming preparation

# 9. Current Project Status

Python ECG processing pipeline completed

Real-time ECG dashboard completed

QRS detection operational

Heart rate estimation operational

Arrhythmia classification operational

Arduino-Python communication verified

Real-time serial architecture prepared

Ready for AD8232 ECG sensor integration

# 10. Future Work

Integrate AD8232 ECG sensor module

Acquire real human ECG signals

Implement live ECG visualization

Optimize real-time filtering

Improve arrhythmia detection accuracy

Add ECG data recording system

Implement wireless monitoring

Develop embedded display system

# 11. Conclusion

A real-time biomedical ECG signal analysis system was
successfully developed using Python and Arduino-based
communication architecture.

The project currently supports:
- Real-time DSP processing
- ECG visualization
- QRS detection
- Heart rate estimation
- Arrhythmia classification
- HIL-ready serial communication

The software platform is now fully prepared for integration
with the AD8232 ECG acquisition hardware module.